Standardisation exercise

The suggested practical exercise described below is not a true standardisation test, however it will help you standardise the way the Hb measurements are taken and select the best Hb measurers. If you have time to conduct a standardisation test, refer to the following publication: Burger S and Pierre-Louis J. A procedure to estimate the accuracy and reliability of HemoCue™ measurements of survey workers. ILSI. 2003

Typically, a training on haemoglobin measurement will contain between 5 to 12 members. For the standardising exercise, each trainee should take two measurements (i.e. filing out two microcuvettes from two different blood drops-blood drop #3 and #4) from two different finger sticks from a minimum of 3 fellow trainees. Use the table below to write down the results and assess the quality of the Hb measurements.

Form for standardiSation exercise

| Volunteer name | Assessing how good the trainee is at filling up the microcuvette |  |  |  |  |  |  |  | Assessing how good the trainee is at finger sticking |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Finger 1 |  |  |  | Finger 2 |  |  |  |  |  |
|  | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 |
|  | Blood drop #3 | Blood drop #4 | C1- C2 | Potential reasons for difference ≥ (+/-) 0.5 g/dL | Blood drop #3 | Blood drop #4 | C5-C6 | Potential reasons for difference≥(+/-) 0.5 g/dL | C1-C5 | Potential reasons for difference ≥ (+/-) 0.5 g/dL |
| 1 |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |

Example of a filled out form

|  | Assessing how good the trainee is at filling up the microcuvette |  |  |  |  |  |  |  | Assessing how good the trainee is at finger sticking |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Finger 1 |  |  |  | Finger 2 |  |  |  |  |  |
| Volunteer name | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 |
|  | Blood drop #3 | Blood drop #4 | C1- C2 | Potential reasons for difference ≥ (+/-) 0.5 g/dL | Blood drop #3 | Blood drop #4 | C5-C6 | Potential reasons for difference≥(+/-) 0.5 g/dL | C1-C5 | Potential reasons for difference ≥ (+/-) 0.5 g/dL |
| 1 | 9.4 | 9.7 | -0.3 | - | 9.7 | 9.3 | 0.2 | - | -0.3 | - |
| 2 | 11.0 | 11.6 | -0.6 | 2nd microcuvette not completely filled | 11.3 | 11.0 | 0.3 | - | -0.3 | - |
| 3 | 10.9 | 12.2 | -1.3 | 2nd microcuvette not completely filled | 11.6 | 11.9 | 0.3 | - | 0.7 | Squeezed finger 1 while filling microcuvette |
| 4 | 12.6 | 12.5 | 0.1 | - | 11.8 | 12.4 | -0.6 | Air bubbles in 1st microcuvette | 0.8 | Air bubbles in finger 2 microcuvette |
| 5 | 10.0 | 12.8 | -2.8 | Alcohol not dry before filling 1st microcuvette | 13.3 | 13.0 | 0.3 | - | -3.3 | Alcohol not dry before filling microcuvette on finger 1 |